## Slide 1: DLS Innovation
Naniwa
February 2024
Nathaniel Dziuba

## Slide 2: Contents
Innovation goals.
What is DLS and how does it work.
Integra Comparison: Manual vs Automated (Colorimetric Test)
DLS Data Analysis
DLS Runs
a. Initial integra run for feasibility and optimal standard (e.g., beads) dilution range

## Slide 3: Innovation Development Goals
Develop a method to determine optimal concentration range of samples.
Automate sample pipetting and serial dilution with Integra Assist Plus.
Automate data analysis to screen sample dilutions.

## Slide 4: Innovation Development Goals: Continued
The purpose of this development is to create a process that can be used on any sample with an unknown concentration. This method could be used to determine the optimal sample concentration range for DLS.
The Integra Assist Plus platform is used to automate sample preparation and plate preparation.
A template generator is used to assist in programming the DLS software for analysis.
Output DLS data is screened using a python script to determine which dilutions generate a near linear relationship in normalized intensity measurements.
The user will be required to take the screened sample list and confirm the integrity of the data by review of the auto-correlation function.

## Slide 5: DLS Background
Particles in solution scatter light
In DLS the intensity of scattered light is measured as a function of time
Variations in scattered light are due to Brownian motion of particles
Intensity vs time plots are analyzed with an autocorrelation function

## Slide 6: DLS Background
Auto-Correlation Function:
  Measures the degree of similarity in time series data
  Data at t=0 is compared incrementally to data farther away.
    Data close to the start of the time series are similar (high correlation) and as we increase in time away from the start the correlation decreases
The auto-correlation function is then fit with different models:
  Cumulant fits are for monodisperse distributions
  Regularization fits are for polydisperse distributions and multimodal species

## Slide 7: DLS: Required Sample Conditions
Scattered light is proportional to the weight-average molar mass and the molecule concentration
LS ~ Mw·c
For this to be true a few assumptions are made:
1) There is enough material to scatter light
2) The analyte has a size smaller than the wavelength of light used for scattering.
    If it isn’t then the Mie scattering occurs.
3) The concentration of analyte does not produce multi-scattering
    Scattered light is scattered again off other particles leading to larger observed average intensities
Sample Criteria:
    Nearly clear
    No particulates or sedimentation (centrifuge before use)
    Operate with known concentrations or perform serial dilution experiment
    Rule-of-thumb is 1-10 mg/mL as a place to start

## Slide 8: Integra Assist Plus

## Slide 9: Integra Manual Vs Automatic
The development goal was to generate a method that could perform a 2-fold serial dilution 7-times in triplicate on a 384 well plate. This would allow for a total of 16 samples on a full plate or 8 samples for half a plate.
The integra assist plus was run multiple times with a simple colometric reagents (i.e., red food colring, SoloVPE calibrant) to determine if the robotic platform functioned correctly.
Red food coloring was used initially for the first few development runs but was later replaced with SoloVPE calibrant due to the dye coating the pipette tips during serial dilution.
A combination of trial and error along with white papers from the vendor were used to determine the proper settings.

## Slide 10: 384-Well Plate Diagram For 2-Fold Dilution Series
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
A | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
B | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
C | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
D | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
E | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
F | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
G | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
H | 1:1 | 1:2 | 1:4 | 1:8 | 1:16 | 1:32 | 1:64 | 1:128
 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
A | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
B | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
C | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
D | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
E | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
F | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
G | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent
H | 44 µL Sample | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent | 22 µL Diluent

## Slide 11: Integra Assist Plus
Experimental Parameters for Integra Assist Plus that were used form comparison to manual
Integra 08 indicates experiment 8 with integra, experiment 9 was manual.
Parameter | Integra 08 | Integra 10
Asp. Disp. Speed | 7 | 6
Mix Vol. | 20 | 15
Liquid Entry/Exit | Normal | Slow
Mix Speed | 7 | 6
SD Source Target | 4.1 mm | 4.8 mm
Pre-Wetting SD | No | No
Air Bubbles Observed? | Yes | No
Blow Out | Yes | Yes

## Slide 12: Integra vs Manual Pipetting Using Dye: Linearity
Chart: Blanked Absorbance vs Dilution Factor
| 1 | 0.5 | 0.25 | 0.125 | 6.25E-2 | 3.125E-2 | 1.5625E-2 | 7.8125E-3
Manual Average | 3.5385 | 1.7669 | 0.74856 | 0.36228 | 0.18367 | 0.085238 | 0.029825 | 0.010888
Integra 10 Average | 3.3575 | 1.8101 | 0.7354 | 0.2926 | 0.12383 | 0.048438 | 0.020037 | 0.0061625
Integra 08 Average | 3.4141 | 1.729 | 0.76293 | 0.33109 | 0.14854 | 0.06625 | 0.026063 | 0.0097875

## Slide 13: Calculated Dilution Factor: Bar Chart with CI Error Bars
Chart: Calculated Dilution Factor From Previous Dilution Error Bars are ± CI
Manual 09 Average | 2.0083 | 2.3798 | 2.0876 | 1.9781 | 2.1603 | 2.8589 | 2.8432
Integra 10 Average | 1.8559 | 2.4636 | 2.522 | 2.3625 | 2.597 | 2.4777 | 3.0649
Integra 08 Average | 1.9753 | 2.2684 | 2.3094 | 2.2381 | 2.2544 | 2.5493 | 2.8101

## Slide 14: Calculated Dilution Factor Statistical Significance
Chart: Calculated Dilution Factor From Previous Dilution Error Bars are ± CI
Manual 09 Average | 2.0083 | 2.3798 | 2.0876 | 1.9781 | 2.1603 | 2.8589 | 2.8432
Integra 08 Average | 1.9753 | 2.2684 | 2.3094 | 2.2381 | 2.2544 | 2.5493 | 2.8101
Dilution step 4 and 6 show significant difference in mean calculated dilution factor in a paired t-test, p<0.05.
Step 6 & 7 are near the limit of the detector.
Step 2 mean + CI = 2.6.
Will use as the limit in data processing.

## Slide 15: DLS Data Analysis Flow Chart

## Slide 16: Running Sample With Plate Templates in Dynamics: Flow Chart

## Slide 17: Dynamics Software Example: Sample
Prepopulate samples 1-12 in the software for future use.
Sample parameters should be reviewed prior to running the experiment.
If the wrong information is input, the data is reprocessed, no reruns required.

## Slide 18: Dynamics Software Example: Plate Template

## Slide 19: Template Example
Samples can be loaded in a CSV file in the software to assign a well with various parameters: Well position, Sample Name, Dilution Factor, Sample.

## Slide 20: DLS Data Processing

## Slide 21: Data Processing

## Slide 22: DLS Data Analysis
Initialization of the data frame (think excel) with basic column header filtering

## Slide 23: DLS Data Analysis
Each measurement is filtered for Baseline, SOS, Amplitude, and # of marked acquisitions
The data is then grouped by Sample Name and Dilution Factor for Normalized Intensity

## Slide 24: DLS Data Analysis
- Results are filtered for a Dilution Factor fold-difference of 2 and for the normalized intensity mean fold-difference to equal 2 ± 30%

## Slide 25: Excel Export
Data that fit the criteria described above are exported into an excel workbook within the working directory.
  All data presented in the